Introduction to Metrology and Characterization Equipment

Objective:  This lab provides an introduction to the operation the ellipsometer, profilometer and optical microscope.  Take detailed notes on the theory and operation behind each tool, while noting the proper use and limitations of each tool.  These tools will be used repeatedly throughout the semester.

Background:  Having the ability to create features on the micro and nanoscale is only useful if there is some way to verify your results.  Therefore, characterization is an integral part of the fabrication process.  A wide variety of characterization tools and techniques are used to ensure reproducible results in every manufacturing process. The nature of the sample and the information to be extracted from the sample dictate which techniques are used. This lab explores three common characterization tools/techniques for analyzing thin films: ellipsometry, profilometry, and optical microscopy. Each technique has its advantages and limitations which are discussed in detail below.

Figure I: Gaertner Scientific L115S-8 Ellipsometer

Ellipsometer:  The Gaertner Scientific L115S-8 Ellipsometer (Figure I) is capable of measuring the index of refraction and material thickness for transparent thin films (0-6000nm) deposited on a reflective substrate.  Many types of thin films can be analyzed with this instrument; polysilicon, different types of oxides and nitrides, etc.  There are some types of materials that cannot be analyzed: organic films such as photoresist, small features, and opaque films like metals.  All films to be analyzed have to be partially transparent.

Ellipsometry is based on the measurement of the light polarization change upon reflection from a sample surface or interface.  This very sensitive measurement technique provides unequalled capabilities for thin film metrology.  As a non-contact optical technique, spectroscopic ellipsometry is non-destructive and offers excellent accuracy.  This includes film thickness for single layer or complex multilayer stacks, ranging from a few angstroms to several microns.  Other film properties can be determined, like the index of refraction of a material.

The shape of the reflected ellipse is measured and film thickness is determined based on information from software integrated within the instrument.  A linearly polarized input beam from the source is converted to an elliptically polarized reflected beam, which is analyzed by the detector to give a readout of the index of refraction, angle of incidence, and film thickness.

Figure II: Veeco Dektak 6M Profilometer

Profilometer:  The Veeco Dektak 6M surface profiler (Figure II) is capable of performing the following functions:

Surface Roughness:  Using software, the tool is capable of measuring the average roughness of the surface of the sample.

Feature Height:  One common mode of operation involves using the stylus to measure the height of features in patterned thin films.

Film Stress:  Based on sample curvature, the profilometer can be used to measure the stress (tensile or compressive) in solid thin films.

The profilometer is capable of measuring a large feature range (1Å to 6.5µm), making it a versatile tool. Since the stylus is in contact with the sample surface, it is important that the profilometer only be used to measure relatively rigid films.  The stylus will damage soft films and cause a distorted image.  The stylus on the profilometer scans across the sample surface in the y-direction.  As it moves across the surface, the stylus moves up and down tracking surface features.  This vertical displacement is recorded by the tool for measurement purposes.

Optical Microscope:  The Leitz Ergolux optical microscope is a reflective microscope capable of viewing samples in both bright and dark field mode.

Bright Field:  A traditional light source is used to illuminate the sample and the reflection yields the image.

Dark Field: light scattered off irregularities on the sample surface making features visible through the microscope.

An optical microscope can be used to image many different samples with relatively large feature sizes. Light reflects off the surface of the sample, allowing an optical image to be obtained through a series of lenses.  An interchangeable objective lens allows the magnification of the microscope to be altered.

Experiment:  This lab will be divided into three stations: the profilometer, ellipsometer, and optical microscope. The group will be divided into sub-groups.  The subgroups will rotate through each station until all of the samples are measured. The goal of this experiment is to learn how to use each tool in a proficient manner.

The film thickness of five samples will be measured at the ellipsometer station.  Samples should be measured using the standard five point measurement system illustrated in Figure III.  The first two measurements on each sample will be two angle measurements and the remaining three will be single angle measurements. Record the measurements in Table I. The samples are labeled 1-5 according to the list below.

- 2100 Å of SiO2 on a silicon substrate.

- 5-20 Å of “native” SiO2 on a silicon substrate.

- 1000 Å of SiO2 on a silicon substrate.

- 450 Ǻ of SiO2 on a silicon substrate.

- 1800 Ǻ of SiNx on a silicon substrate.

Feature heights and widths of three patterned samples will be measured at the profilometer station. Record the measurements in Table II.  The samples are labeled 1-3 according to the list below.

- 1.4µm of patterned photoresist (PR) on a SiO2 substrate.

- Remaining PR + SiO2 etch depth of post etch sample.

- Etch depth of SiO2 post etch/post PR removal.

One sample will be examined at the microscope station. Each subgroup should practice taking bright and dark field images. Bright field images should contain measurements made with the imaging software. Be sure to save one bright field and one dark field image to be included in the formal lab report.

Figure III.  Wafer map for five point measurement.

| Ellipsometer | Measured Thickness (Å) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample | 2 angle |  | 2 angle |  | specific | specific | specific | Average |
|  | 1 |  | 2 |  | 3 | 4 | 5 |  |
| 1 | 50 |  | 50 |  |  |  |  |  |
|  | 70 |  | 70 |  |  |  |  |  |
| 2 | 50 |  | 50 |  |  |  |  |  |
|  | 70 |  | 70 |  |  |  |  |  |
| 3 | 50 |  | 50 |  |  |  |  |  |
|  | 70 |  | 70 |  |  |  |  |  |
| 4 | 50 |  | 50 |  |  |  |  |  |
|  | 70 |  | 70 |  |  |  |  |  |
| 5 | 50 |  | 50 |  |  |  |  |  |
|  | 70 |  | 70 |  |  |  |  |  |

Table I:  Film thickness measurement results for the ellipsometer.

| Profilometer | Include Unit of Features Measured |  |
| --- | --- | --- |
| Sample | Width | Height |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |

Table II:  Feature width and height measurement results for the profilometer.

Questions:

- For a pattern of photoresist on SiO2, which tool from this lab, would you select to measure:

  - Width of areas to be etched?  Why?  (3)

  - Depth of photoresist?  Why?  (3)

- For each tool used in this lab discuss the principle of operation.  Be sure to discuss whether the tool functions in a destructive or a non-destructive manner (1)

- Name and discuss two limitations of the ellipsometer. (2)

- Name and discuss two limitations of the profilometer. (2)

- Explain a limitation of the microscope for measuring feature dimensions. (3)

- Explain when bright field and dark field should be used on the microscope.  (2)

- Complete Table I from the lab handout (3)

- Complete Table II from the lab handout (3)

- Include a bright field image with measurements at 100X (3)

- Include a dark field image with measurements at 10X (3)

- For the profilometer, how did the feature height and width change from scan to scan and what factors influenced the change in feature dimensions (3)

Lab questions are due by 11:45 AM the Monday following lab completion.

No work will be accepted late!