## Slide 1: Chemical sensor arrays
Microcantilevers Learning Module

## Slide 2: Unit Overview
Chemical Sensor Arrays (CSA) are MEMS (Microelectromechanical Systems) used to gather, detect, measure, and identify a substance or several substances in a minute sample.
CSAs can be cantilever-based, using micro and nano-sized cantilevers.
This unit is an overview of the operation and applications of MEMS Chemical Sensor Arrays.
Notes: That sample could be a drop of blood, an unknown gas, or an unknown liquid.

## Slide 3: Objectives
State at least five applications of a MEMS CSA.
Compare and contrast the operations of a static and dynamic cantilever-based CSA.
Describe at least two operating characteristics of a CSA.

## Slide 4: Introduction
A CSA is …
an array of microtransducers and supporting integrated circuits.
designed to detect and measure the amount or concentration of one or more substances in a sample environment.
The substances (target materials or analytes) could be specific gas molecules or atoms, antibodies or proteins.
Process flow of a Chemical Sensor Array

## Slide 5: MEMS Chemical Sensor Arrays
The micro-size makes a CSA attractive for applications from biomedical to aerospace.
CSAs can analyze blood samples, detect fuel leaks, analyze a scent or dissect a taste.
Most MEMS CSAs consist of an array of microcantilever transducers.
Microcantilever Array

## Slide 6: CSA Applications
  Medical diagnostics
  Forensics
  Environmental control
  Fragrance design
  Food production
  Security and defense
  Detection of chemical vapors (see graphic right)
  Detection of biological agents
Optical Array for detection of
Chemical vapors

## Slide 7: Lab-On-a-Chip (LOC) Array
LOC is a MEMS that incorporates several functions on a single chip.
Chemical sensing
Pumping microfluids
Electronics
Lab-on-a-chip – Blood Analysis Chemical Sensor Array
[Photo courtesy of Y. Tai, California Institute of Technology]
A miniaturized, portable blood-count device used by astronauts to analyze blood samples, in real-time, to diagnose infection, allergies, anemia or deficiencies in the immune system.

## Slide 8: CSA as Olfactory Sensor (Enose)
Artificial olfactory system or nose (Enose)
Analyzes a fragrance by separating the particles that provide an overall scent.
Used to detect specific compounds in a food's odor.
Determine the freshness of the meat or the presence of contaminants.
The picture is the ENose developed by NASA and tested on the International Space Station in late 2008. The ENose Sensor Unit (the darker-looking metal object) is housed in its Interface Unit (white). It is currently used in the space station to detect ammonia.
[Graphic source: NASA]

## Slide 9: How Does the ENose Work?
A collection of 16 different polymer films on a set of electrodes. (The graphic (a) illustrates six films/electrodes.)
Films are specially designed to conduct electricity based on the film’s resistance.
A baseline resistance reading is established with no odors (ambient air).

## Slide 10: How Does the ENose Work?
When a substance -- such as the stray molecules from an ammonia leak -- is absorbed into these films, the films expand slightly (b), changing their resistivity.
The amount of expansion of each film determines the amount of its electrode current.

## Slide 11: How Does the ENose Work?
Each polymer reacts differently to a chemical compound.
Changes in resistivity in a single polymer film would not be enough to identify a compound; however, the cumulative changes of the 16 films produce a distinctive, identifiable pattern for a specific compound.
This graphic shows a different compound being sensed.

## Slide 12: Biosensors
Analyze samples for substances such as antibodies, proteins, antigens, and DNA.
Glucose monitors
pH sensors
DNA or gene microarrays analyze and measure the activity of genes.
Gene expression values from microarray experiments can be represented as heat maps to visualize the results of data analysis. The green represents reduced gene expression or activity.
[Image is public domain. Image source: Wikipedia: Gene Expression Profiling]
Notes: DNA microarrays enable scientists and doctors to analyze complex biological problems

## Slide 13: Challenging Environments
Due to their micro size, design and packaging, MEMS CSAs are used in environments that are destructive to comparable macrosensors and where other types of sensors are ineffective.
  Electric and magnetic fields
  Hazardous chemical vapors
  Nuclear radiation
  Radio frequency (RF) radiation
  Contaminated and hazardous liquids

## Slide 14: Cantilever-Based Sensor Array
The most common CSA transducer is the microcantilever.
Its versatility and low construction costs make it an ideal transducer for a variety of analytes.
Microcantilevers (see SEM picture) are typically 10 – 500 μm long, up to 100 μm wide, and up to 2 μm thick. A pitch of 250 mm is standard in optical-fiber array applications.
Scanning Electron Microscope (SEM) image of a Cantilever Sensor Array (
These cantilevers were developed by the Swiss Nanoscience Institute for proteomic and genomic applications.
[Image courtesy of Dr. Christoph Gerber, Institute of Physics, University of Basel]

## Slide 15: How does a Cantilever-Based CSA work?

## Slide 16: The Microcantilever of a CSA
Microcantilever transducers.
Suspended end is coated with a chemically sensitive probe coating which experiences a chemical change when an analyte is adsorbed.
Different probe coatings on each cantilever can allow for the detection of several different substances within the same sample.
In the above graphic, the probe coating is deposited on a gold layer. Each coating is designed to detect one and only one target material (analyte). The back microcantilever is the reference transducer.

## Slide 17: CSA Block Diagram
Analyte binds to the probe coating causing a minute, but measurable change in the cantilever's mechanical or electrical properties.
The change is measured and processed by the integrated circuitry (Signal Processing).
The data is analyzed and compared to reference data for determining the type and amount of analytes (Analytes Identification).

## Slide 18: An Interesting Fact
Prior to the onset of micron and nano-technology, these minute changes in the mechanical properties of such small devices were considered negligible.
Current technology provides innovative methods for measuring these negligible changes allowing microscopic components to be monitored and measured like their macroscopic equivalents.

## Slide 19: Mass-Sensitive Transducers
Microcantilever transducers have a sensitivity compared to similar macrotransducers.
Overall mass is measurably affected by the chemisorption of a very small quantity or mass of material in the probe coating.
A small change in mass causes a measurable change in one or more of the cantilever's properties.
Mass-Sensitive Transducer

## Slide 20: Stress-Induced Curvature
Mass-sensitive transducers become less effective as the transducer and analytes get smaller.
For nanosized analytes, cantilever displacement is caused by surface stress in the chemisorption of the analytes with the probe coating.
The adsorption the coating to expand or contract which bends the cantilever.
This is "stress-induced curvature."

## Slide 21: Static and Dynamic CSAs
Static and Dynamic modes of operation are used in CSAs.
The static mode measures the bending of the cantilever.
This bending is a static response, but it is measurable.
The dynamic mode measures a shift in the cantilever's resonant frequency.

## Slide 22: Cantilever Construction
For both the static and dynamic modes, the cantilevers are constructed with different surface coatings.
The coating may cover only the tip of the suspended or the entire surface (functionalizing the surface).
Selective coating deposition process is referred to as "functionalizing the surface.
Functionalizing the Surface – Selecting coating a portion of the cantilever surface with the probe coating.

## Slide 23: Static Mode
As the chemically reactive surface selectively adsorbs the analytes, the cantilever bends due to surface stress.
Bending causes a measurable cantilever displacement.
Displacement is measured by sensing a change in angular deflection or resistance.
Functionalizing the Surface – Selecting coating a portion of the cantilever surface with the probe coating.

## Slide 24: Static - Measuring ∆R (resistance)
Piezoresistive material is fabricated into the microcantilever.
As the target material is adsorbed by the chemically reactive layer the cantilever bends.
This creates a measurable change in the resistance of the piezoresistive material.
Amount of change is representative of the amount of analytes or their concentration within the sample.

## Slide 25: Static – Measuring ∆Angular Deflection
A reflective layer is coated onto the surface of the microcantilever prior to the chemically reactive layer.
A laser beam is reflected off the cantilever's surface creating a reference angle.
As the cantilever bends, the change in the angular deflection is detected by measuring the change in position of the reflected beam.

## Slide 26: Dynamic Mode
Microcantilevers are initially excited by external actuation.
Excitation results in an oscillation at resonant frequency.
Analytes adsorb to the probe coating.
Adsorption changes the cantilever's mass resulting in a change in resonant frequency.
Amount of change in resonant frequency depends upon the concentration of the analytes and the amount of exposure time.

## Slide 27: Static or Dynamic?
In liquid environments the damping effect of the liquid on the cantilever's movement can result in false readings.
Therefore, static CSA's are primarily used in liquid environments.
Gaseous environments use both static and dynamic CSAs.

## Slide 28: Operating Characteristics of CSAs
Operating characteristics include
  Sensitivity
  Selectivity
  Response time
  Size
  Power consumption
Cantilever-based CSAs have proven to be
highly sensitive,
highly selective, and with
fast response times.

## Slide 29: Mass Sensitivity
Microcantilever transducers have an inherently high mass sensitivity due to their small mass.
The physical properties (width, thickness, length and material) are used to further enhance its sensitivity.
Which cantilever would have a higher resonant frequency
    short or long?
    thick or thin?

## Slide 30: Response Time
Response time for a microcantilever is the time it takes for the cantilever to respond to the target material on its surface and produce a change in the output.
Response time is affected by several parameters, three of which are the
  concentration of the target material in the environment,
  the probe material itself, and
  the method used to interpret the change in a mechanical property.

## Slide 31: Other Types of CSAs
The type of CSA primarily discussed in this unit was a cantilever-based CSA that uses mass or stress sensitive transducers.
There are variations of CSAs that use other types of transducers that may be better suited for a specific application and its requirements.

## Slide 32: Optical Sensor Arrays
A chemical reaction between the probe material and the analytes affects an optical property of the transducer such as color (wavelength) or light intensity.
The graphic shows a partial output from a colorimetric optical array which acts as an "optoelectronic nose“.
It uses an array of multiple dyes whose color changes are based on the full range of intermolecular interactions with the sample analytes and transducers.
The four volatile organic compounds in the graphic have four different patterns as identified by the sensor array.

## Slide 33: Schottky Diode Sensor Array
The target substance absorbs (diffuses) through the probe material to a metal layer.
The metal layer serves as the gate for a diode.
Any change at the gate causes a change in the diode's characteristics.
Here is a link to an animation on how a Schottky Diode Sensor works:
http://www.nsf.gov/od/lpa/news/press/01/pr01105flash.htm
Schottky Diode Sensor Array
(University of California, Santa Barbara, Department of Chemical Engineering.]

## Slide 34: Cell-based Sensor Arrays
Cell-based Sensors use biological cells as the transducers to detect the presence of specific molecules within the cell's environment.
One type of molecular transducer uses cell amplification.
When a cell interacts with the analytes a chemical change occurs causing the production of many "so-called second messenger" molecules.
This is a biological gain that can be measured to determine the amount of analytes in the sample.

## Slide 35: CSAs Working Together
With the variety of sensor arrays available, a system can be developed to mimic the human senses.
  Cantilever-based arrays distinguish between different smells and tastes,
  Optical arrays react to different wavelengths and intensities, and
  Acoustic arrays detect a change in acoustic properties as a result of interacting with the environment.

## Slide 36: Food For Thought
What are some additional applications where CSA's could be used to detect a combination of gases, scents or particles?
In the dynamic mode, which microcantilever would be more sensitive to mass loading – one 100 microns in length or 60 microns in length? (Assume the thickness, width and materials are the same for both cantilevers).

## Slide 37: Food For Thought
In many applications, dynamic mode or static mode CSAs could be used. This is a CSA used to detect a specific virus in the bloodstream. Based on your knowledge of the microcantilever modes of operation, which mode – dynamic or static – do you think would be best for this application and why?
Nanocantilevers coated with antibodies* (blue-green) that capture viruses (red spheres). As the cantilevers identify and capture more virus molecules, one or more of the mechanical or electrical characteristics of the cantilevers can change and be detected by an electronic interface. The size of the particle being detected and captured is one of the factors affecting the size of the cantilever. [Image generated and printed with permission by Seyet, LLC]
*Antibodies are proteins produced in the blood in response to the presence of an antigen (e.g., virus, bacteria, toxin).

## Slide 38: Summary
A cantilever-based CSA uses an array of microcantilevers to detect and measure specific materials within a sample environment.
The micro size of the cantilevers results in higher selectivity, improved sensitivity, faster response time and low construction costs.
These characteristics make the cantilever CSA a very popular sensor for a wide range of applications.

## Slide 39: Acknowledgements
Made possible through grants from the National Science Foundation Department of Undergraduate Education #0830384, 0902411, and 1205138.
Any opinions, findings and conclusions or recommendations expressed in this material are those of the authors and creators, and do not necessarily reflect the views of the National Science Foundation.
Southwest Center for Microsystems Education (SCME) NSF ATE Center
© 2012 Regents of the University of New Mexico
Content is protected by the CC Attribution Non-Commercial Share Alike license.
Website: www.scme-nm.org
Revised Feb 2017